Standardization of your Flow Cytometry Experiment from Day to Day in DiVa

Day 1

- Set up your experiment settings through your controls. This generally takes into account the positive and negative populations to ensure you are in the most sensitive range of the cytometer.

- Run samples

- Add an extra tube in a separate specimen labeled "Bead Control"

- Create an extra worksheet labeled "Daily Beads”

- Create Histograms for every channel utilized in your experiment

- Run beads* and set interval gates around positive and negative peaks (*acceptable beads include single mid-peak beads, or CST, or 6 or 8 peak rainbow beads where you use the middle peak- if you’re not including the negative beads)

- Save application settings to a location on your shared drive

- Duplicate experiment without data and delete all but the two specimens with one tube each.

- Create a template from duplicated/truncated experiment (Right click experiment and export->experiment template)

Day X

- Create new experiment from your template (Experiment->New Experiment, find your template)

- Accept CST settings

- Apply application settings (right click Cytometer settings and apply)

- Run bead control and verify peaks hit interval gates (and if they don’t, change voltages until they do)

- Run samples